## Start
Englischsprachige Fassung
AG-ME Eichbehörden der Länder
- Arbeitsausschuss Waagen -
Verification of Compatibility for Modules of
Weighing Instruments
|  | Printing: | Data,
Evaluation for
|  |  | - | One range weighing instruments
|  |  | - | Two ranges weighing instruments
|  |  | - | Three ranges weighing instruments
|  |  | - | Two-interval weighing instruments
|  |  | - | Three-interval weighing instruments
|  | State: | 26th of November 2004
Copyright by AA-Waagen
X00A-805-130_R320

## Data-input
Eichbehörden der Länder, englischsprachige Fassung
Manufacturer:
a) Metrological and technical Data : |  |  | TRUE
Weighing instrument (WI) | accuracy class |  | class |  |  |  | -
Scale number: | maximum capacity | verification scale interval | Max | e |  |  | kg | kg
XXX | [6 empty] | kg | kg
Type of weighing instrument: | [6 empty] | kg | kg
YYY | [6 empty] | kg | kg
| reduction ratio |  | R |  |  |  | -
| number of load cells |  | N |  |  |  | -
| initial zero setting range |  | IZSR |  |  |  | kg
TAC (Type Approval Certificate): | correction for non uniform distributed load |  | NUD |  |  |  | kg
T12994 | dead load of load receptor |  | DL |  |  |  | kg
| additive tare |  | T+ |  |  |  | kg
| limits of temperature range |  | Tmin | Tmax | -20 | 40 | ° C | ° C
Owner of TAC: | length of cable |  | L |  |  |  | m
RINSTRUM Pty Ltd. | cross section of wire |  | A |  |  |  | mm2
Auswertegerät (AWG)
Electronic indicator | accuracy class |  | class |  | III |  | -
Manufacturer: | max. number verification scale intervals |  | nind |  | 4000 |  | -
RINSTRUM Pty Ltd. | load cell excitation voltage |  | Uexc |  | 5 |  | V
| minumum inputvoltage for the indicator |  | Umin |  | 0 |  | mV
Type: | min. input voltage per verification scale interval |  | Dumin |  | 0.8 |  | mV
R320/xxxx (ser.-No.) | minimum/maximum load cell resistance |  | RLmin | RLmax | 43 | 3500 | W | W
TC (Test certificate): | limits of temperature range |  | Tmin | Tmax | -20 | 40 | ° C | ° C
TC11437 | fraction of the maximum permissible error |  | pind |  | 0.5 |  | -
and/or TAC (Type Approval Certificate): | cable connection (4- or 6 wire) |  | number of wires |  |  |  | -
| max. value of cable length per wire cross section |  | (L/A)max |  |  |  | m/mm2
Load cell(s) (LC) | accuracy class |  | class |  |  |  | -
Manufacturer : | maximum capacity |  | Emax |  |  |  | kg
| minimum dead load |  | Emin |  |  |  | kg
| rated output |  | C |  |  |  | mV/V
Type: | max. number verification scale intervals |  | nLC |  |  |  | -
| minimum verification scale interval or |  | vmin |  |  |  | kg
| the ratio Y = E max / v min |  | Y |  |  |  | -
TC (Test certificate): | ratio Z = E max / (2 DR) or |  | Z |  |  |  | -
| minimum dead load output return |  | DR |  |  |  | kg
| input resistance of single load cell |  | RLC |  |  |  | W
or TAC (Type Approval Certificate): | limits of temperature range |  | Tmin | Tmax |  |  | ° C | ° C
| fraction of the maximum permissible error |  | pLC |  | 0.7 |  | -
Verbindungselemente
Connecting elements | fraction of the maximum permissible error |  | pcon |  |  |  | -
The above described modules have been applied in originally Version. Date and signature of an authorized representative of the WI-Manufacturer:
GM-P9, Anhang 8.1.10 zu Abschnitt 9.1 |  | 2004-11-26 | [5 empty] | page 1 of 6
Data input |  | X00A-805-130_R320
Note on G10: Accuracy class: The accuracy classes of NAWI, Ind and LCs are compatible, if they follow table below: NAWI I II III IIII Ind I I*), II II*), III III, IIII LC A A*), B B*), C C, D *) at sufficient temperature range and evaluation of stability against humidity and creep corresponds to requirement in the lower (more accurate) class
Note on G11: Maximum capacity of One range weighing instrument
Note on G12: Maximum capacity of the lower weighing range of multiple range or multi-interval weighing instrument
Note on G13: Maximum capacity of the 2nd weighing range of multipe range or multi-interval weighing instruments (Max1 = lower weighing range)
Note on G14: Maximum capacity of the upper weighing range of multiple range or multi-interval weighing instruments (Max1 = lower weighing range)
Note on G15: reduction ratio R The reduction ratio is given by a lever work system. According to R76 No T.3.3 R = (Force onto the load cell)/(Gravitation force onto the load receptor) Using lever works normally the number N of load cells is N = 1 and R < 1 The value of R has to be given by the WI-manufacturer Without lever system the reduction ratio R always will be R = 1 Format of Field: Number Excample of wrong format 1:4, inputs with ":" will be treated as time correct format 1/4 or 0,25
Note on G16: number of load cells
Note on G17: initial zero setting range: maximal range, which is automatically pulled to zero by switching on the weighing instrument, before indicating. If change of load receptor is possible by the owner for retail scales 20% of Max (when realizable with indicator) in other cases 4% of Max
Note on G18: factor NUD: (WI = weighing instrument, LC = load cell) Factor for Non Uniform Distributed load caused by excentric loading, NUD shall prohibit that a load cell is applied with load outside of the garanteed measuring range. According to WELMEC 2 (Issue 3) Nr. 3.1.6.6 followig values may be accepted for NUD: - WIs with lever work and with one LC, WIs with one single point LC and WIs with load receptors which allöw only minimal eccentric load application (eg hopper scales, but without shaker on the receptor ) 0 % Max - other conventional instruments : 20 % Max - over head track scales, fork lift scales and weighbridges : 50 % Max - Multi-platform WIs with more than one load receptor fix combined 50 % Max (total) variable selection or combined 50 % Max (of single platform)
Note on G19: dead load: weight of the load receptor and additionally mounted constructions on the load receptor
Note on G20: maximum value of the additive tare device (if existing) if no device of Additive Tare, fill out with 0
Note on G21: limit of the temperature range lower limit of the environment temperature for the whole weighing instrument
Note on G22: length of cable length of cable between indicator and load cell junction box ( load cell cables themselves have to maintain original)
Note on G23: cross section of wire cross section of one wire from cable (4 or 6 wires) between indicator and load cell junction box Note - By application of different cross sections in one cable use for 4-wires the supply wires for 6-wires the reference wires - Using flash protection elements, the exitation voltage at the load cell has to be checked when applying check on condition (8)
Note on G25: Accuracy class: The accuracy classes of NAWI, Ind and LCs are compatible, if they follow table below: NAWI I II III IIII Ind I I*), II II*), III III, IIII LC A A*), B B*), C C, D *) at sufficient temperature range and evaluation of stability against humidity and creep correspond to requirement in the lower (more accurate) class
Note on G26: maximal number of verification scale intervals for the indicator: The maximal number n ind of verification scale intervals of the indicator must be equal or better than the number of verification scale intervals of the weighing instrument n = Max/e. For multiple range and muti-interval weighing instruments this must be met for every weighing range.
Note on G27: power supply for load cell(s) Note: Insert total voltage, wrong + 5 V correct 10 V Further see comment for cross section A
Note on G28: the indicator may need a minimum input voltage at all (bias), if so the value is given in the indicator documentation; this indicator is not able to measure changing polarity of voltage If no Umin is required insert 0
Note on G29: the indicator needs a minimum voltage per verification scale interval, which is given in the indicator documentation
Note on G30: lower limit of the impedance of the load cell(s) to be supplied the load cells are connected in parallel (limitted by power supply fittted by the indicator)
Note on G31: limit of the temperature range lower limit of the environment temperature for the indicator
Note on G32: fraction of the maximum permissible error of the weighing instrument used for the indicator according to R76 No 3.5.4 standard pind = 0,5 ; allowable 0,3 > pind > 0,8
Note on G33: cable connection (4- or 6-wire) between load cell junction box and indicator The original 4-wire load cell cable have to be kept in any case, The connection of original 4-wire load cell cables just beside the indicator is also allowed in any case
Note on G34: Maximal allowed ratio of cable length per wire cross section see comment to cross section of wire A
Note on G36: Accuracy class: The accuracy classes of NAWI, Ind and LCs are compatible, if they follow table below: NAWI I II III IIII Ind I I*), II II*), III III, IIII LC A A*), B B*), C C, D *) at sufficient temperature range and evaluation of stability against humidity and creep correspond to requirement in the lower (more accurate) class
Note on G37: maximum capacity of one load cell The maximum capacity of the load cell shall not be exeeded when applying Max on the weighing instrument.
Note on G38: minimum dead load of one load cell the applied dead load on the load cell in the WI must be bigger or equal to the minimum dead load of the load cell when Emin = 0 insert 0
Note on G39: rated output of the load cell when loading with maximum capacity (in mV/V)
Note on G40: maximal number of verification scale intervals of the load cell: The max. number nLC of verification scale intervals of the load cell must be equal or better than the number of verification scale intervals of the weighing instrument n = Max/e. For multiple range and muti-interval weighing instruments this must be met for every weighing range.
Note on G41: smallest verification interval of one load cell If Y or vmin is not indicated (f.e. in the test certificate), nLC will be used instead of Y .
Note on G42: ratio of Emax / vmin of the load cell If Y or vmin is not indicated (f.e. in the test certificate), nLC will be used instead of Y .
Note on G43: ratio of Emax / 2* DR of the load cell Where DR or Z not known (f.e. not indicated in the test certificate), nLC will be used instead of Z.
Note on G44: minimum dead load return of the load cell: Where DR or Z not known (f.e. not indicated in the test certificate), nLC will be used instead of Z.
Note on G45: Input impedance of one load cell: The WI may have N load cells. The smallest ratio of RLC / N is limitted by the power supply of the indicator
Note on G46: limit of the temperature range lower limit of the environment temperature for the load cell
Note on G47: fraction of the maximum permissible error of the weighing instrument applied for the load cells if not indicated in the test certificate , pLC = 0,7.
Note on G49: fraction of the maximum permissible error of the weighing instrument applied for connecting elements pcon , For the fraction of all unknown mechanical and electrical connection elements between load receptor, load cell(s) a d indicator. acceptable according to R76 No 3.4.5 pcon = 0,5
Note on H11: verification interval of One range weighing instrument
Note on H12: verification interval of the lower weighing range of a multiple range or multi-interval weighing instrument
Note on H13: verification interval of the 2nd weighing range of multiple range or multi-interval weighing instrument (e 1 = smallest weighing range)
Note on H14: verification interval of the upper weighing range of multi-interval weighing instrument (e1 = smallest verification interval)
Note on H21: limit of the temperature range upper limit of the environment temperature for the whole weighing instrument
Note on H30: upper limit of the impedance for the load cell(s) to be supplied (limitted by sensitivity against EMC)
Note on H31: limit of the temperature range upper limit of the environment temperature for the indicator
Note on H46: limit of the temperature range upper limit of the environment temperature for the load cell

## One range WI
Eichbehörden der Länder, englischsprachige Fassung
b) Verification of Compatibility - One range weighing instrument (WI)
(1) Accuracy class of weighing instrument (WI) compatible to class of indicator (ind) and load cell (LC) | [6 empty] | I. O. ?
| 5 |  | 3 |  |  | 0
| LC | & | IND | equal or better |  | WI
|  | & | III | equal or better |  |  | NO
(2) Temp.limits of the weighing instr.(WI) compared with the temp.limits of the load cell (LC) and the indicator (IND)
|  | LC |  | IND |  | WI
| Tmin |  | & | -20 | £ | -20 | NO
| Tmax |  | & | 40 | ³ | 40 | NO
(3) Sum of the squares of the fractions pi of the max. permissible errors of connecting elements, indicator and load cells
| pcon2 | + | pind2 | + | pLC2 | £ 1
|  | + | 0.25 | + | 0.49 | £ 1 | NO
(4) Number of verification scale intervals of the weighing instrument and the indicator
|  |  |  | nind | ³ | n=Max/e
|  | One range weighing instrument |  | 4000 | ³ |  | NO
| Mehrteilungs- |  | i = 1 | 0 | ³ | 0
| oder Mehrbereichswaage |  | i = 2 | 0 | ³ | 0
|  |  | i = 3 | 0 | ³ | 0
(5) Maximum capacity of load cells must be compatible to Max of the weighing instrument
Factor Q (EN 45 501 No 4.12.1): Q = (Max+DL+IZSR+NUD+T+)/Max = | [5 empty] | DL wanted
|  |  | (Q*Max*R)/N |  | £ | Emax
|  |  | R wanted |  | £ |  | NO
(6a) Maximum number of verification scale intervals of load cell and number of scale intervals of the weighing instrument
|  |  |  | nLC | ³ | n=Max/e
|  | One range weighing instrument |  |  | ³ |  | NO
| Mehrteilungs- |  | i = 1 | 0 | ³ | 0
| oder Mehrbereichswaage |  | i = 2 | 0 | ³ | 0
|  |  | i = 3 | 0 | ³ | 0
(6b) Rückkehr des Vorlastsignals der WZ und kleinster Eichwert e1 einer Mehrteilungswaage
|  |  | n LC oder Z=Emax/(2*DR) |  | ³ | Maxr / e1
|  |  | 0 |  | ³ | 0
(6c) Rückkehr des Vorlastsignals der WZ und kleinster Eichwert e1 einer Mehrbereichswaage
|  |  | nLC oder Z=Emax/(2*DR) |  | ³ | 0,4*Maxr/e1
|  |  | 0 |  | ³ | 0
(6d) Minimum dead load of the load cells to the actual dead load of the load receptor
|  |  |  | DL*R/N | ³ | Emin
[5 empty] | ³ |  | NO
(7) Verification scale interval of the weighing instrument and minimum load cell scale interval must be compatible
|  |  |  | e*R/ ÖN | ³ | vmin= Emax/Y
[5 empty] | ³ |  | NO
(8) Minimum input voltage for the indicator, minimum input voltage per verification scale interval and actual output of the LCs
|  | minimum input voltage | U= C*Uexc*R*DL/(Emax *N) |  | ³ | Umin
|  | (unloaded WI) |  |  | ³ | 0 | NO
|  | input voltage per verification interval | Du = C*Uexc*R*e/(Emax *N) |  | ³ | Dumin
[5 empty] | ³ | 0.8 | NO
(9) Allowed impedance range for the electronic indicator and actual load cell impedance
|  | RLmin | £ | RLC / N | £ | RLmax
|  | 43 | £ | R LC wanted | £ | 3500 | NO
(10) Cable length per wire cross section of the connection cable between the load cell(s) and indicator
|  |  |  | (L/A) | £ | (L/A)max
[5 empty] | £ |  | NO
GM-P9, Anhang 8.1.10 zu Abschnitt 9.1 |  |  |  | 2004-11-26 |  |  | page 2 of 6
Evaluation One range weighing instrument |  |  |  | X00A-805-130_R320

## Two ranges WI (hidden)
Eichbehörden der Länder, englischsprachige Fassung
b) Verification of Compatibility - Two ranges weighing instrument (WI)
(1) Accuracy class of weighing instrument (WI) compatible to class of indicator (IND) and load cell (LC) | [6 empty] | I. O. ?
| 5 |  | 3 |  |  | 0
| LC | & | IND | equal or better |  | WI
|  | & | III | equal or better |  |  | NO
(2) Temp.limits of the weighing instr.(WI) compared with the temp.limits of the load cell (LC) and the indicator (IND)
|  | LC |  | IND |  | WI
| Tmin |  | & | -20 | £ | -20 | NO
| Tmax |  | & | 40 | ³ | 40 | NO
(3) Sum of the squares of the fractions pi of the max. permissible errors of connecting elements, indicator and load cells
| pcon2 | + | pind2 | + | pLC2 | £ 1
|  | + | 0.25 | + | 0.49 | £ 1 | NO
(4) Number of verification scale intervals of the weighing instrument and the indicator
|  |  |  | nind | ³ | n(i) = Max(i) /e(i)
| Einbereichswaage |  |  | 4000 | ³ | 0
| two ranges weighing instrument |  | i = 1 | 4000 | ³ |  | NO
|  |  | i = 2 | 4000 | ³ |  | NO
|  |  | i = 3 | 0 | ³ | 0
(5) Maximum capacity of load cells must be compatible to Max of the weighing instrument
Faktor Q (EN 45 501 No 4.12.1) : Q = (Max2+DL+IZSR+NUD+T+)/Max2 = | [5 empty] | DL wanted
|  |  | (Q*Max2*R)/N |  | £ | Emax
|  |  | R wanted |  | £ |  | NO
(6a) Maximum number of verification scale intervals of load cell and number of scale intervals of the weighing instrument
|  |  |  | nLC | ³ | n(i) = Max(i) /e(i)
| Einbereichswaage |  |  | 0 | ³ | 0
| two ranges weighing instrument |  | i = 1 |  | ³ |  | NO
|  |  | i = 2 |  | ³ |  | NO
|  |  | i = 3 | 0 | ³ | 0
(6b) Rückkehr des Vorlastsignals der WZ und kleinster Eichwert e1 einer Mehrteilungswaage
|  |  | n LC oder Z=Emax/(2*DR) |  | ³ | Maxr / e1
|  |  | 0 |  | ³ | -0.5
(6c) Minimum dead load output return of the load cell and smallest verification scale interval e1 of a multiple range weighing instrument
|  |  | nLC oder Z= Emax / (2*DR) |  | ³ | 0,4*Max2 /e1
|  |  | 0 |  | ³ |  | NO
(6d) Minimum dead load of the load cells to the actual dead load of the load receptor
|  |  |  | DL*R/N | ³ | Emin
[5 empty] | ³ |  | NO
(7) Verification scale interval of the weighing instrument and minimum load cell scale interval must be compatible
|  |  |  | e1*R/ÖN | ³ | vmin= Emax /Y
[5 empty] | ³ |  | NO
(8) Minimum input voltage for the indicator, minimum input voltage per verification scale interval and actual output of the LCs
|  | minimum input voltage | U= C*Uexc*R*DL/(Emax *N) |  | ³ | Umin
|  | (unloaded WI) |  |  | ³ | 0 | NO
|  | input voltage per verification interval | Du = C*Uexc*R*e/(Emax *N) |  | ³ | Dumin
[5 empty] | ³ | 0.8 | NO
(9) Allowed impedance range for the electronic indicator and actual load cell impedance
|  | RLmin | £ | RLC / N | £ | RLmax
|  | 43 | £ | R LC wanted | £ | 3500 | NO
(10) Cable length per wire cross section of the connection cable between the load cell(s) and indicator
|  |  |  | (L/A) | £ | (L/A)max
[5 empty] | £ |  | NO
GM-P9, Anhang 8.1.10 zu Abschnitt 9.1 |  |  |  | 2003-01-21 |  |  | page 3 of 6
Evaluation Two ranges weighing instrument

## Three ranges WI (hidden)
Eichbehörden der Länder, englischsprachige Fassung
b) Verification of Compatibility - Three ranges weighing instrument (WI)
(1) Accuracy class of weighing instrument (WI) compatible to class of indicator (IND) and load cell (LC) | [6 empty] | I. O. ?
| 5 |  | 3 |  |  | 0
| LC | & | IND | equal or better |  | WI
|  | & | III | equal or better |  |  | NO
(2) Temp.limits of the weighing instr.(WI) compared with the temp.limits of the load cell (LC) and the indicator (IND)
|  | LC |  | IND |  | WI
| Tmin |  | & | -20 | £ | -20 | NO
| Tmax |  | & | 40 | ³ | 40 | NO
(3)Sum of the squares of the fractions pi of the max. permissible errors of connecting elements, indicator and load cells
| pcon2 | + | pind2 | + | pLC2 | £ 1
|  | + | 0.25 | + | 0.49 | £ 1 | NO
(4) Number of verification scale intervals of the weighing instrument and the indicator
|  |  |  | nind | ³ | n(i) = Max(i) /e(i)
| Mehrbereichswaage |  |  | 4000 | ³ | 0
| three ranges weighing instrument |  | i = 1 | 4000 | ³ |  | NO
|  |  | i = 2 | 4000 | ³ |  | NO
|  |  | i = 3 | 4000 | ³ |  | NO
(5) Maximum capacity of load cells must be compatible to Max of the weighing instrument
Faktor Q (EN 45 501 No 4.12.1) : Q = (Max3+DL+IZSR+NUD+T+)/Max3 = | [5 empty] | DL wanted
|  |  | (Q*Max3*R)/N |  | £ | Emax
|  |  | R wanted |  | £ |  | NO
(6a) Maximum number of verification scale intervals of load cell and number of scale intervals of the weighing instrument
|  |  |  | nLC | ³ | n(i) = Max(i) /e(i)
| Einbereichswaage |  |  | 0 | ³ | 0
| three ranges weighing instrument |  | i = 1 |  | ³ |  | NO
|  |  | i = 2 |  | ³ |  | NO
|  |  | i = 3 |  | ³ |  | NO
(6b) Rückkehr des Vorlastsignals der WZ und kleinster Eichwert e1 einer Mehrteilungswaage
|  |  | n LC oder Z=Emax/(2*DR) |  | ³ | Maxr / e1
|  |  | 0 |  | ³ | -0.5
(6c) Minimum dead load output return of the load cell and smallest verification scale interval e1 of a multiple range weighing instrument
|  |  | nLC oder Z=Emax/(2*DR) |  | ³ | 0,4*Max3 /e1
|  |  | 0 |  | ³ |  | NO
(6d) Minimum dead load of the load cells to the actual dead load of the load receptor
|  |  |  | DL*R/N | ³ | Emin
[5 empty] | ³ |  | NO
(7) Verification scale interval of the weighing instrument and minimum load cell scale interval must be compatible
|  |  |  | e1*R/ÖN | ³ | vmin= Emax/Y
[5 empty] | ³ |  | NO
(8) Minimum input voltage for the indicator, minimum input voltage per verification scale interval and actual output of the LCs
|  | minimum input voltage | U= C*Uexc*R*DL/(Emax *N) |  | ³ | Umin
|  | (unloaded WI) |  |  | ³ | 0 | NO
|  | input voltage per verification interval | Du = C*Uexc*R*e/(Emax *N) |  | ³ | Dumin
[5 empty] | ³ | 0.8 | NO
(9) Allowed impedance range for the electronic indicator and actual load cell impedance
|  | RLmin | £ | RLC / N | £ | RLmax
|  | 43 | £ | R LC wanted | £ | 3500 | NO
(10) Cable length per wire cross section of the connection cable between the load cell(s) and indicator
|  |  |  | (L/A) | £ | (L/A)max
[5 empty] | £ |  | NO
GM-P9, Anhang 8.1.10 zu Abschnitt 9.1 |  |  |  | 2003-01-21 |  |  | page 4 of 6
Evaluation Three ranges weighing instrument

## Two-interval WI (hidden)
Eichbehörden der Länder, englischsprachige Fassung
b) Verification of Compatibility - Two-interval weighing instrument (WI)
(1) Accuracy class of weighing instrument (WI) compatible to class of indicator (IND) and load cell (LC) | [6 empty] | I. O. ?
| 5 |  | 3 |  |  | 0
| LC | & | IND | equal or better |  | WI
|  | & | III | equal or better |  |  | NO
(2) Temp.limits of the weighing instr.(WI) compared with the temp.limits of the load cell (LC) and the indicator (IND)
|  | LC |  | IND |  | WI
| Tmin |  | & | -20 | £ | -20 | NO
| Tmax |  | & | 40 | ³ | 40 | NO
(3)Sum of the squares of the fractions pi of the max. permissible errors of connecting elements, indicator and load cells
| pcon2 | + | pind2 | + | pLC2 | £ 1
|  | + | 0.25 | + | 0.49 | £ 1 | NO
(4) Number of verification scale intervals of the weighing instrument and the indicator
|  |  |  | nind | ³ | n(i) = Max(i) /e(i)
| Einbereichswaage |  |  | 4000 | ³ | 0
| two-interval weighing instrument |  | i = 1 | 4000 | ³ |  | NO
|  |  | i = 2 | 4000 | ³ |  | NO
|  |  | i = 3 | 0 | ³ | 0
(5) Maximum capacity of load cells must be compatible to Max of the weighing instrument
Faktor Q (EN 45 501 No 4.12.1) : Q = (Max2+DL+IZSR+NUD+T+)/Max2 = | [5 empty] | DL wanted
|  |  | (Q*Max2*R)/N |  | £ | Emax
|  |  | R wanted |  | £ |  | NO
(6a) Maximum number of verification scale intervals of load cell and number of scale intervals of the weighing instrument
|  |  |  | nLC | ³ | n(i) = Max(i) /e(i)
| Einbereichswaage |  |  | 0 | ³ | 0
| two-interval weighing instrument |  | i = 1 |  | ³ |  | NO
|  |  | i = 2 |  | ³ |  | NO
|  |  | i = 3 | 0 | ³ | 0
(6b) Minimum dead load output return of the load cell and smallest verification scale interval e1 of a multi-interval weighing instrument
|  |  | nLC oder Z=Emax/(2*DR) |  | ³ | Max2 / e1
|  |  | 0 |  | ³ |  | NO
(6b) Rückkehr des Vorlastsignals der WZ und kleinster Eichwert e1 einer Mehrbereichswaage
|  |  | nLC oder Z=Emax/(2*DR) |  | ³ | 0,4*Maxr/e1
|  |  | 0 |  | ³ | 0
(6d) Minimum dead load of the load cells to the actual dead load of the load receptor
|  |  |  | DL*R/N | ³ | Emin
[5 empty] | ³ |  | NO
(7) Verification scale interval of the weighing instrument and minimum load cell scale interval must be compatible
|  |  |  | e1* R /ÖN | ³ | vmin= Emax/Y
[5 empty] | ³ |  | NO
(8) Minimum input voltage for the indicator, minimum input voltage per verification scale interval and actual output of the LCs
|  | minimum input voltage | U = C*Uexc*R*DL/(Emax *N) |  | ³ | Umin
|  | (unloaded WI) |  |  | ³ | 0 | NO
|  | input voltage per verification interval | Du = C*Uexc*R*e/(Emax *N) |  | ³ | Dumin
[5 empty] | ³ | 0.8 | NO
(9) Allowed impedance range for the electronic indicator and actual load cell impedance
|  | RLmin | £ | RLC / N | £ | RLmax
|  | 43 | £ | R LC wanted | £ | 3500 | NO
(10) Cable length per wire cross section of the connection cable between the load cell(s) and indicator
|  |  |  | (L/A) | £ | (L/A)max
[5 empty] | £ |  | NO
GM-P9, Anhang 8.1.10 zu Abschnitt 9.1 |  |  |  | 2003-01-21 |  |  | page 5 of 6
Evaluation Two-interval weighing instrument

## Three-interval WI (hidden)
Eichbehörden der Länder, englischsprachige Fassung
b) Verification of Compatibility - Three-interval weighing instrument (WI)
(1) Accuracy class of weighing instrument (WI) compatible to class of indicator (IND) and load cell (LC) | [6 empty] | I. O. ?
| 5 |  | 3 |  |  | 0
| LC | & | IND | equal or better |  | WI
|  | & | III | equal or better |  |  | NO
(2) Temp.limits of the weighing instr.(WI) compared with the temp.limits of the load cell (LC) and the indicator (IND)
|  | LC |  | IND |  | WI
| Tmin |  | & | -20 | £ | -20 | NO
| Tmax |  | & | 40 | ³ | 40 | NO
(3) Sum of the squares of the fractions pi of the max. permissible errors of connecting elements, indicator and load cells
| pcon2 | + | pind2 | + | pLC2 | £ 1
|  | + | 0.25 | + | 0.49 | £ 1 | NO
(4) Number of verification scale intervals of the weighing instrument and the indicator
|  |  |  | nind | ³ | n(i) = Max(i) /e(i)
| Einbereichswaage |  |  | 4000 | ³ | 0
| three-interval weighing instrument |  | i = 1 | 4000 | ³ |  | NO
|  |  | i = 2 | 4000 | ³ |  | NO
|  |  | i = 3 | 4000 | ³ |  | NO
(5) Maximum capacity of load cells must be compatible to Max of the weighing instrument
Faktor Q (EN 45 501 No 4.12.1) : Q = (Max3+DL+IZSR+NUD+T+)/Max3 = | [5 empty] | DL wanted
|  |  | (Q*Max3*R)/N |  | £ | Emax
|  |  | R wanted |  | £ |  | NO
(6a) Maximum number of verification scale intervals of load cell and number of scale intervals of the weighing instrument
|  |  |  | nLC | ³ | n(i) = Max(i) /e(i)
| Einbereichswaage |  |  | 0 | ³ | 0
| three-interval weighing instrument |  | i = 1 |  | ³ |  | NO
|  |  | i = 2 |  | ³ |  | NO
|  |  | i = 3 |  | ³ |  | NO
(6b) Minimum dead load output return of the load cell and smallest verification scale interval e1 of a multi-interval weighing instrument
|  |  | nLC oder Z=Emax/(2*DR) |  | ³ | Max3 / e1
|  |  | 0 |  | ³ |  | NO
(6c) Rückkehr des Vorlastsignals der WZ und kleinster Eichwert e1 einer Mehrbereichswaage
|  |  | nLC oder Z=Emax/(2*DR) |  | ³ | 0,4*Maxr/e1
|  |  | 0 |  | ³ | 0
(6d) Minimum dead load of the load cells to the actual dead load of the load receptor
|  |  |  | DL*R/N | ³ | Emin
[5 empty] | ³ |  | NO
(7) Verification scale interval of the weighing instrument and minimum load cell scale interval must be compatible
|  |  |  | e1*R/ÖN | ³ | vmin= Emax/Y
[5 empty] | ³ |  | NO
(8) Minimum input voltage for the indicator, minimum input voltage per verification scale interval and actual output of the LCs
|  | minimum input voltage | U= C*Uexc*R*DL/(Emax *N) |  | ³ | Umin
|  | (unloaded WI) |  |  | ³ | 0 | NO
|  | input voltage per verification interval | Du = C*Uexc*R*e/(Emax *N) |  | ³ | Dumin
[5 empty] | ³ | 0.8 | NO
(9) Allowed impedance range for the electronic indicator and actual load cell impedance
|  | RLmin | £ | RLC / N | £ | RLmax
|  | 43 | £ | R LC wanted | £ | 3500 | NO
(10) Cable length per wire cross section of the connection cable between the load cell(s) and indicator
|  |  |  | (L/A) | £ | (L/A)max
[5 empty] | £ |  | NO
GM-P9, Anhang 8.1.10 zu Abschnitt 9.1 |  |  |  | 2025-04-01 |  |  | page 6 of 6
Evaluation Three-interval weighing instrument

## Tabelle1 (hidden)
One range WI
Two ranges WI
Three ranges WI
Two-interval WI
Three -interval WI
1